## Slide 1: Ultrasound Research Platform for Transmission of Coded Excitation Signals
Emma Muir, Sam Muir, Jacob Sandlund, & David Smith
Advisor: Dr. José Sánchez
Co-Advisor: Dr. James Irwin

## Slide 2: Ultrasound Imaging

## Slide 3: Quantitative Ultrasound
[1]
Benign
Malignant

## Slide 4: Introduction Outline
Introduction
How Ultrasound Works
Coded Excitation
Objective
Motivation
Significance
Design Comparison

## Slide 5: How Ultrasound Works

## Slide 6: Coded Excitation
Conventional Ultrasound [2]
Coded Excitation Ultrasound [2]

## Slide 7: Coded Excitation Platforms
Research Platforms
Mostly single-element
Large multi-element
  RASMUS
RASMUS [3]

## Slide 8: Objective
Ultrasound Research Platform Prototype
  Arbitrary Waveforms
    Coded excitation signals
  Multi-element
    Beamforming
  Reduced size and cost
Lecroy Oscilloscope

## Slide 9: Motivation
Improve…
  Ultrasound Techniques
  Ultrasound Research
Reduce size and cost

## Slide 10: Significance
Medical Applications
  Detect and Diagnose Tumors
  Noninvasive
  Faster Results

## Slide 11: Design Comparison
Previous Designs:
Our Design:
Digital Device
Amplifier
Transducer
Digital Device
Transducer
Switching Amplifier
D/A

## Slide 12: Our Method
Oversample
  1-bit
  Densities represent
  voltages

## Slide 13: Our Method
Transducer acts as a (BP) filter
  Smooths / Averages

## Slide 14: Encoding Differences
Example:
  0.5 V DC
  -1 to 1 V Dynamic range
8-bit Two’s Complement (-128 to 127):
  Value = 64 (0100 0000)
Sigma Delta Modulation:
  Oversample
  1-bit

## Slide 15: Outline
Introduction
Functional Description
Methods
Results and Discussion
Conclusion
Questions

## Slide 16: Outline
Introduction
Functional Description
Methods
Results and Discussion
Conclusion
Questions

## Slide 17: System Requirements
Up to 4 transducer channels
Excitations <= 3 μs
SNR > 50 dB

## Slide 18: System Block Diagram
18

## Slide 19: System Block Diagram
19
Generate Waveform

## Slide 20: System Block Diagram
20
Transmit Waveform

## Slide 21: System Block Diagram
21
Receive Image

## Slide 22: System Block Diagram
22
Create Image

## Slide 23: PC Data Processing
23

## Slide 24: PC Data Processing
24

## Slide 25: PC Data Processing
25

## Slide 26: PC Data Processing
26

## Slide 27: PC Data Processing
27

## Slide 28: Data Processing Methods: Conversion to an Image
Time Gain Compensation (TGC)
  Attenuation
  TGC = Att * Depth * (Probe frequency)
  white noise for larger depths

## Slide 29: PC Data Processing
29

## Slide 30: Data Processing Methods: Conversion to an Image
Envelope Detection
  Determines the bounds of the processed signal
  Detects width and contains the display information
  Absolute value of the Hilbert Transform

## Slide 31: PC Data Processing
31

## Slide 32: Outline
Introduction
Functional Description
Methods
Results and Discussion
Conclusion
Questions

## Slide 33: Methods
Sigma Delta Modulation
PC/FPGA Interface
FPGA
Data Processing
  Pulse Compression
  Delay Sum Beamforming

## Slide 34: System Block Diagram
34

## Slide 35: Sigma-Delta Modulation Requirements
< 10% Mean Squared Error (MSE)
500 M samples/second
Accuracy vs. Overloading (Saturation)
  Order = 2nd
  OSR = 16
    must be a power of 2
    16*2 = 32 samples per period

## Slide 36: Sigma-Delta Modulation Methods
36
[4]

## Slide 37: Sigma-Delta Modulation Methods
[4]

## Slide 38: Sigma-Delta Modulation Methods
38

## Slide 39: Sigma-Delta Modulation Methods
[5]

## Slide 40: System Block Diagram
40

## Slide 41: PC/FPGA Interface Requirements
Assign waveform to pins
  Independent for each pin
  (3 μs) * (500 MHz) = 1500 bits/waveform
  1500 + 36 = 1536 bits/waveform (divisible by 512)
Assign delay to pins
  Increments of 4ns = (1/250 MHz)
  250 MHz = memory clock rate of FPGA

## Slide 42: PC/FPGA Interface Requirements
Transfer information for 4 pins in < 1 sec
  <32 sec for 128 pins
  (4 pins) * (1536 bits/waveform) sent within 1 sec
  ~6 Kbps
Start transmission

## Slide 43: PC/FPGA Interface Methods
UART connection
  115200 baud
    Fastest FPGA baud rate
  Sends as
    1 start bit
    8 data bits
    2 stop bits
  (1536/8)*11*4 = 8448 bits
  ~73 ms for 4 channels
  ~2.3 s for 128 channels
Start
Stop
Data

## Slide 44: PC/FPGA Interface Methods

## Slide 45: FPGA Requirements
Transmit at 500 MHz
Output waveforms in parallel
  4 individualized waveforms
  Length of 3 ms per waveform
  1536-bits per waveform

## Slide 46: FPGA Methods
Yes
No

## Slide 47: FPGA Methods
Transmit at 500 MHz
  Two 250 MHz clock edges (transmits on rising and falling edge)
  250 MHz * 2 = 500 MHz
XOR

## Slide 48: System Block Diagram
48

## Slide 49: Data Processing Requirements
Data Processing
  Less than 2 minutes
Display an image
  Depths between 0.25 cm and 30 cm
  Dynamic range between 40 dB and 60 dB

## Slide 50: PC Data Processing
50

## Slide 51: Data Processing Methods: Pulse Compression
Restore the spatial resolution
Match reflected wave to original excitation
Use Wiener filter
  Optimal solution between a match filter and an inverse filter [6]
  Solution determined by
    Smoothing Factor (SF)
    Predicted signal-to-noise-ratio (SNR)
Predict SNR = 50 dB

## Slide 52: Data Processing Methods: Pulse Compression
Matched filter
  Cross correlation of original coded excitation and received signal
  Creates side lobes
  Does not amplify noise
  Optimal for large noise – small SNR
Inverse filter
  Inverse of the original coded excitation
  No side lobes
  Amplifies noise
  Optimal for no noise – large SNR

## Slide 53: Data Processing Methods: Pulse Compression
Wiener Filter Equation
= Coded Excitation
= Smoothing Factor
= SNR of system
Noise increases
SNR decreases
λ/S increases
Closer to a Match Filter
Noise decreases
SNR increases
λ/S decreases
Closer to an Inverse Filter
[1]

## Slide 54: Data Processing Methods: Pulse Compression
SNR = 60 dB

## Slide 55: PC Data Processing
55

## Slide 56: Data Processing Methods: Beamforming
128 Sensor Array Transducer
Focal Point
Focal Point
Narrowest beam
Greatest amplitude
Beamforming not
necessary at this point
20 mm
4 mm
38.36 mm

## Slide 57: Data Processing Methods: Beamforming
Amplitude at Point = Σi=1 Si( Depth + Delay(Si,Point)) [7]
Delay(S,P) = (DSP - DSF)/c [7]
S = sensor P = point
DSP = distance from sensor to point
DSF = distance from sensor to point
c = 1540m/s (speed of sound in tissue)
8
Sensors
Focal Point
Point
Depth

## Slide 58: Outline
Introduction
Functional Description
Methods
Results and Discussion
Conclusion
Questions

## Slide 59: Experiment
59

## Slide 60: Experiment
60

## Slide 61: Chirp Transmission
61

## Slide 62: Experiment
62

## Slide 63: Transducer Model
63

## Slide 64: Experiment
64

## Slide 65: Chirp Reconstruction
65
1.34% MSE

## Slide 66: Chirp Reconstruction
66

## Slide 67: Correlation Results
Correlations | Filtered Sigma-delta Modulated Linear Chirp | Filtered Captured Data
Filtered Linear Chirp | 99.84% | 99.33%
Filtered Sigma-delta Modulated Linear Chirp | -------- | 99.53%

## Slide 68: Correlation Results
Correlations | Filtered Sigma-delta Modulated Linear Chirp | Filtered Captured Data
Filtered Linear Chirp | 99.84% | 99.33%
Filtered Sigma-delta Modulated Linear Chirp | -------- | 99.53%

## Slide 69: Correlation Results
Correlations | Filtered Sigma-delta Modulated Linear Chirp | Filtered Captured Data
Filtered Linear Chirp | 99.84% | 99.33%
Filtered Sigma-delta Modulated Linear Chirp | -------- | 99.53%

## Slide 70: Correlation Results
Correlations | Filtered Sigma-delta Modulated Linear Chirp | Filtered Captured Data
Filtered Linear Chirp | 99.84% | 99.33%
Filtered Sigma-delta Modulated Linear Chirp | -------- | 99.53%

## Slide 71: Experiment
71

## Slide 72: Pulse Compression
72

## Slide 73: REC Chirp
73
h1(n) * c1(n) = h2(n) * c2(n)

## Slide 74: Data Processing Simulations
74
Field II Software [8]
10 mm separation
46 dB SNR
10
20
30
40
50
60
70
80
90
10
20
-10
-20
0
Distance in mm
100
Transducer

## Slide 75: Beamforming Results
Without Beamforming
With Beamforming

## Slide 76: REC Chirp Simulation
76
REC Excitation and Pulse Compression
Impulse Excitation

## Slide 77: Graphical User Interface
77

## Slide 78: Outline
Introduction
Functional Description
Methods
Results and Discussion
Conclusion
Questions

## Slide 79: Conclusion
Valid waveform transmission
Portable system
Multi-channel
Research potential
Lecroy Oscilloscope

## Slide 80: Acknowledgements
The authors would like to thank Analog Devices and Texas instruments for their donation of parts.
This work is partially supported by a grant from Bradley University (13 26 154 REC)
Dr. Lu
Mr. Mattus
Mr. Schmidt
Andy Fouts

## Slide 81: References
[1] J. R. Sanchez et al., "A Novel Coded Excitation Scheme to Improve Spatial and Contrast Resolution of Quantitative Ultrasound Imaging," IEEE Trans. Ultrason. Ferroelectr. Freq. Control, vol. 56, no. 10, pp. 2111-2123, October 2009.
[2] "Clinical Image Library." GE Healthcare-. GE Healthcare. Web. 14 Apr. 2011. <http://www.gehealthcare.com/usen/ultrasound/ products/msul7im.html>
[3] J. A. Jensen et al., “Ultrasound Research Scanner for Real-time Synthetic Aperture Data Acquisition,” IEEE Trans. Ultrason., Ferroelect., Freq. Contr., vol. 52, no. 5, pp. 881–891, 2005.
[4] R. Schreier and G. C. Temes. Understanding Delta-Sigma Data Converters, John Wiley & Sons, Inc., 2005.
[5] R. Schreier, The Delta-Sigma Toolbox Version 7.3. Analog Devices, Inc, 2009.

## Slide 82: References Cont.
[6] T. Misaridis and J. A. Jensen, “Use of Modulated Excitation Signals in Medical Ultrasound Part I: Basic Concepts and Expected Benefits,” IEEE Trans. Ultrason. Ferroelectr. Freq. Control, vol. 52, no. 2, pp. 177-191, February 2005.
[7] Thomeniu, Kai E. "Evolution of Ultrasound Beamformers." IEEE Ultrasonics Symposium (1996): 1615-622. Print.
[8] J.A. Jensen. Field: A Program for Simulating Ultrasound Systems, Medical & Biological Engineering & Computing, pp. 351-353, Volume 34, Supplement 1, Part 1, 1996.

## Slide 83: Questions?